Competitive Enzyme Immunoassay Kit for

Quantitative Analysis of Furaltadone metabolite (AMOZ)

1. Background

Nitrofurans are synthetic broad-spectrum antibiotics, which are frequently employed in animal product for its excellent antibacterial and pharmacokinetic properties. They had been also used as growth promoters in pig, poultry and aquatic product. In long term studies with experimental animals the parent drugs and their metabolites showed carcinogenic and mutagenic characteristics. This has led to a prohibition of nitrofurans for the treatment of animals used for food product. The nitrofuran drugs furaltadone, nitrofurantoin and nitrofurazone were banned in food animal product in the EU in 1993, and the use of furazolidone was prohibited in 1995.

The analysis of residues of nitrofuran drugs needs to be based on the detection of the tissue bound metabolites of the nitrofuran parent drugs. Since the parent drugs are very rapidly metabolized, and the tissue bound nitrofuran metabolites will retain for a long time, therefore the metabolites are used as the target in the detection of the abuse of nitrofurans. Furazolidone metabolite (AOZ), Furaltadone metabolite (AMOZ), Nitrofurantoin metabolite (AHD) and Nitrofurazone metabolite (SEM).

AMOZ-residues are determined most commonly with LC-UV, LC-MS, or LC-MS/MS techniques. Enzyme immunoassays, compared with chromatographic methods, show considerable advantages regarding sensitivity, detection limit, technical equipment and time requirement.

2. Test Principle

This ELISA kit is designed to detect furaltadone metabolite (AMOZ) based on the principle of “indirect-competitive” enzyme immunoassay. The microtiter wells are coated with capture BSA-linked antigen. AMOZ in sample competes with the antigen coated on the microtiter plate for the antibody. After the addition of horseradish peroxidase labeled anti-antibody, tetramethyl-benzidine (TMB)/peroxide substrate is used and the signal is measured by spectrophotometer. The absorption is inversely proportional to the AMOZ concentration in the sample.

3. Applications

This kit can be used in quantitative and qualitative analysis of AMOZ residue in animal tissue (chicken, pig, cattle, fish and shrimp), milk and honey.

4. Cross-reactions

Furaltadone metabolite (AMOZ)…………………… 100%

furazolidone metabolite…………………………< 0.1%

nitrofurantoin metabolite………………………<0.1%

Furacilinum metabolite…………<0.1%

furazolidone………………………………<1%

Furaltadone………………………………11.1%

nitrofurantoin……………………………<1%

Furacilinum………………………………<1%

5. Materials Required

5.1 Equipments:

┅┅Microtiter plate spectrophotometer (450nm/630nm)

┅┅Rotary evaporator or nitrogen drying instruments

┅┅Homogenizer or stamocher

┅┅Shaker

┅┅Gyroscope or vortex mixer

┅┅Centrifuge

┅┅Analytical balance (inductance: 0.01g)

┅┅Graduated pipette: 10ml

┅┅Rubber pipette bulb

┅┅Volumetric flask: 10ml, 100ml, 1L

┅┅Glass flask: 10ml

┅┅Polystyrene centrifuge tube: 2ml, 10ml, 50ml

┅┅Glass centrifuge tube: 10ml, 15ml

┅┅Micropipettes: 20ul~200ul, 200ul~1000ul,

250ul-multipipette

5.2 Reagents:

┅┅Ethyl acetate (AR)

┅┅N-Hexane (or n-heptane) (AR)

┅┅dipotassium hydrogen phosphate trihydrate

(K2HPO4.3H2O )(AR)

┅┅Methyl alcohol (AR)

┅┅Sodium nitroprusside dihydrate (Na2Fe(CN)5·NO·2H2O) (AR, Milk only)

┅┅Zinc sulfate heptahydrate (ZnSO4·7 H2O) (AR, Milk only)

┅┅Concentrated hydrochloric acid

┅┅Sodium hydroxide (NaOH) (AR)

┅┅Deionized water

6. Kit Components

1、Microtiter plate with 96 wells coated with coupling antigen

2、Standard solutions(6 bottles×1ml/bottle)

0ppb，0.1ppb，0.3ppb，0.9ppb，2.7ppb，8.1ppb

3、High concentration standard control (1ml/bottle )  100ppb

4、enzyme-labeled secondary antibody   12ml …… red cap

5、antibody work  solution  7ml……green cap

6、Solution A    7ml ……… white cap

7、Solution B    7ml ……… red cap

8、Stop solution   7ml ………yellow cap

9、20×Concentrated wash solution  40ml…transparent cap

10、2×concentrated extraction solution  50ml……… blue cap

11、2-nitro-Benzaldehyde  15.1mg… ………black cap

7. Reagents Preparation:

Solution 1: Derivative Reagent:

Add methyl alcohol to the bottle with 2-nitrobenzaldehyde and dilute to 10ml. (at the concentration of 10mM)

Solution 2: 0.1M dipotassium hydrogen phosphate solution

Weigh 22.8g dipotassium hydrogen phosphate trihydrate and dilute to 1L with deionized water.

Solution 3: Solution C (Milk only):

0.36M Sodium nitroprusside (Na2Fe(CN)5·NO) solution: Dissolve 12.5g sodium nitroprusside dihydrate (Na2Fe(CN)5·NO·2H2O) with deionized water diluted to 100ml.

Solution D (Milk only):

1M zinc vitriol solution: Dissolve 29.8g Zinc sulfate heptahydrate (ZnSO4·7 H2O) with deionized water diluted to 100ml.

Solution 4: 1M Hydrochloric acid solution: Dilute 8.3ml concentrated hydrochloric acid with deionized water to 100ml.

Solution 5: 1M Sodium hydroxide (NaOH) solution:

4.0g sodium hydroxide (NaOH) dissolved with deionized water diluted to 100ml.

Solution 6: Extraction solution: Dilute the extraction concentrate solution with deionized water in the volume ratio of 1:1, which will be used for sample extraction. The diluted extraction solution can be conserved for a month at 4℃.

Solution 7: Wash solution: Dilute the wash concentrate solution with deionized water in the volume ratio of 1:19, which will be used for washing the plates. The diluted solution can be stored at 4℃ for 1 month.

8. Sample Preparation

8.1 Notice and precautions for the users before operation:

(a) Please use disposable tips in the process of experiment, and change the tips if different reagent is absorbed.

(b) Make sure that all experimental instruments are clean.

(c) Derivative reagent can be stored for 6 months at 2-8℃.

(d) dipotassium hydrogen phosphate solution can be stored for 3 months at 2-8℃.

(e) HCl can be kept at room temperature for three months

(e) NaOH solution can be kept at room temperature for three months

(d) Store the untreated sample in freeze.

(e) The treated sample can be stored for 24 hours, 2-8℃, in darkness.

8.2 Muscle, liver and aquatic products (fish and shrimp):

┅┅Homogenize the samples with homogenizer or stamocher;

┅┅Then follow Process 8.5;

Note: Samples should be stored in cool dark place and kept in freeze.

8.3 Milk

┅┅Centrifuge milk sample: 10min, at least 3000g, at 10℃. Remove the supernatant fat.

┅┅Transfer 5ml milk samples (without fat) into glass centrifuge tube and add 250(l solution C and 250(l solution D (See solution 3).

┅┅Shake completely with shaker.

┅┅Centrifuge: 10min, at least 3000g, at 4~12 ℃(39-54℉). Or centrifuge after cooling the sample to 8℃(46℉) if without constant temperature centrifuge.

┅┅Then follow Process 8.5;

8.4 Honey

┅┅Weigh 1±0.05g honey into a new polystyrene centrifuge tube (50ml)

┅┅Dissolve with 4ml deionized water and shake completely.

┅┅Add 0.5ml 1M hydrochloric acid solution (see solution 4), and 100(l derivative reagent and shake with shaker completely.

┅┅Do from the 2nd step of Process 8.5 labeled with “*”.

8.5 Process followed the above.

┅┅Weigh 1.0±0.05g of the homogenate (tissue and liver samples), or take 1.1ml of the supernatant milk sample(equal to 1ml milk sample), dissolve with 4ml deionized water, shake if necessary until it is dissolved completely, Add 0.5ml 1M hydrochloric acid solution and 100(l derivative reagent and shake completely for 2min .

* ┅┅Incubation at 37 ℃ over night. (Approx. 16 hours)

┅┅Add 0.1M Potassium phosphate dibasic (see solution 2) 5ml, 1M Sodium hydroxide (see solution 5) 0.4ml and Ethyl acetate 5ml separately, shake completely for 30s.

┅┅Centrifuge: 10min, at least 3000g, at 20-25℃ (68-77℉).

┅┅Transfer the ethyl acetate layer 2.5ml into a dry glass centrifuge tube(10ml) and evaporate organic compound completely at 50~60℃with nitrogen stream or rotary evaporator..

┅┅Dissolve the leftover with 1ml n-hexane (or n-heptane) and shake for 30s, and then add 1ml diluted extraction solution (see solution 6) and shake for 1min.

┅┅Centrifuge: 10min, at least 3000g, at 20-25℃ (68-77℉).

┅┅Remove the supernatant, and use 50 (l per well of the substrate for assay.

9. Enzyme immunoassay assay process

9.1 Notice before assay:

9.1.1 Make sure all reagents and microwells are all at room temperature (20-25℃) for more than 30min until rewarm.

9.1.2 Return all the rest reagents to 2～8℃ immediately after used.

9.1.3 Wash the microwells correctly is an important step in the process of assay, it is the vital factor to the reproducibility of the ELISA analysis.

9.1.4 Avoid the light and cover the microwells during incubation.

9.2 Assay process:

9.2.1 Take all reagents out at room temperature (20-25℃) for more than 30min, homogenize before use.

9.2.2 Get the microwells needed out and return the rest into the zip-lock bag at 2-8℃ immediately.

9.2.3 The diluted wash solution should be rewarmed to room temperature before use.

9.2.4 Number: Numbered every microwell positions and all standards and samples should be run in duplicate. Record the standards and samples positions.

9.2.5 Add standard solution/sample: Add 50 µl of each standard solution or prepared sample to separate duplicate wells. Add 50µl antibody solution. Mix gently by shaking the plate manually and incubate for 30min at 25℃ with cover.

9.2.6 Wash: Remove the cover gently and pure the liquid out of the wells and wash the microwells with 250µl diluted wash solution (see solution 7) at intervals of 10s. Tap the microwells holder upside down vigorously against absorbent paper to ensure complete removal of liquid from the wells (repeat 4-5 times).

9.2.7 Add enzyme-labeled secondary antibody: add enzyme-labeled secondary antibody 100(l/ to each well, mix gently, then cover it and incubate at 25℃ in darkness for 30min, after that, take out and repeat the Wash step.

9.2.8 Colour: Add 50µl solution A and 50µl solution B to each well. Mix gently by shaking the plate manually and incubate for 15min at 25℃ with cover (see notice 12.8).

9.2.9 Measure: Add 50µl the stop solution to each well. Mix gently by shaking the plate manually and measure the absorbance at 450nm against an air blank (It’s suggested measure with the dual-wavelength of 450/630nm. Read the result within 5min after addition of stop solution. ) (We can also measure by sight without stop solution in short of the ELIASA instrument)

There are 2 different methods to determinate the results. Method 1 leads to a round estimation, and method 2 leads to definite quantity estimation. (Please notice: the absorption is inversely proportional to the concentration of AMOZ in the sample.)

10.1 Round estimation

We can get the range of different concentration from the compare of average absorption and standards by sight. For example, the absorption of sample 1 is 0.301, sample 2 is 0.830, and the absorptions of AMOZ standard solutions: 2.018(0 ppb); 1.475(0.1ppb); 1.030(0.3ppb); 0.472(0.9 mg/L); 0.187(2.7 ppb); 0.079(8.1 ppb). So we can say the concentration of diluted sample 1 is between 0.9 ppb to 2.7 ppb; and diluted sample 2 between 0.3 ppb to 0.9 ppb. In order to obtain the AMOZ actually contained in a sample, the diluted sample results must be further multiplied by the corresponding dilution factor.

10.2 Definite quantity estimation

The mean values of the absorbance values obtained for the standards and the samples are divided by the absorbance value of the first standard (zero standard) and multiplied by 100%. The zero standard is thus made equal to 100% and the absorbance values are quoted in percentages.

B

Absorbance (%) = —— ×100%

B0

B ——absorbance standard (or sample)

B0 ——absorbance zero standard

---To draw a standard curve: The absorbance value of standards as y-axis, semilogarithmic of the concentration of the standards (ppb) as x-axis.

--- The AMOZ concentration of each sample (ppb), which can be read from the calibration curve, is multiplied by the corresponding dilution rate of each sample followed, and the actual concentration of sample is obtained.

Please notice:

For evaluation of the ELISA kits special software has been developed for exact and rapid analysis. The analysis software can be ordered on request.

Dilution multiple

---aquatic products (fish and shrimp), muscle, liver: 2.

---milk sample: 2.

---honey sample: 2

11. Sensitivity, accuracy and precision

Test Sensitivity: 0.1ppb

Detection limit:

For the tissue sample will interfere with the test result, the detection limit then becomes 0.3ppb.

Milk and honey sample……………………0.2ppb

Accuracy

Tissue(aquatic products, liver, chicken)………  75±15%           honey ……………………………………………… 90±15%

milk…………………………………………………  90±10%

Precision:

Variation coefficient of the ELISA kit is less than 10%.

12. Notice

12.1 The mean values of the absorbance values obtained for the standards and the samples will be reduced if the reagents and samples have not been regulated to room temperature (20-25℃).

12.2 Do not allow microwells to dry between steps to avoid unsuccessful repetitiveness and operate the next step immediately after tap the microwells holder.

12.3 homogenize each reagent before use.

12.4 Keep your skin away from the stop solution for it is the 2M H2SO4 solution.

12.5 Don’t use the kits out of date. Don’t exchange the reagents of different batches, or else it will drop the sensitivity.

12.6 Storage constitution:

Keep the ELISA kits at 2-8℃ without freeze. Avoid straight sunlight during all incubations. Covering the microtiter plates is recommended.

12.7 Indications for the reagents going bad:

Substrate solution should be abandoned if it turns colors.

The reagents may be turn bad if the absorbance value (450/630nm) of the zero standard is less than 0.5(A450nm＜0.5).

12.8 The coloration reaction need 10-15min after the addition of solution A and solution B; But you can prolong the incubation time ranges from 20min to more if the color is too light to be determined. Never exceed 30min, On the contrary, shorten the incubation time properly.

12.9 The best competitive reaction temperature and reaction temperature are both 25℃. Too high or too low temperature will both lead to the changes of sensitivity and absorbance values.

13. Storage condition and storage period

Storage condition: 2-8℃.

Storage period: 6 months.