Competitive Enzyme Immunoassay Kit for

Quantitative Analysis of β-Estradiol

1. Background

β-Estradiol is an anabolic steroid hormone, which is broadly applied in promoting animal growth. It is now being banned or prohibited in developed countries such as US and the EU for obvious carcinogenicity.  Document No. 235 of Ministry of Agriculture prescribes that the residual of β-Estradiol in animal-derived food must not be detected. GC is the common approach to detect this drug, but it is restricted for the complicated operation and high expense. While this kit is a new product for drug residual detection based on ELISA technology, which only costs 1.5 hours in each operation and can considerably minimize operation errors and work intensity

2. Test Principle

This kit is based on indirect-competitive ELISA technology. The microtiter wells are coated with coupling antigen. β-Estradiol residual in the sample competes with the antigen coated on the microtiter plate for the antibody. After the addition of horseradish peroxidase labeled anti-antibody, tetramethyl-benzidine (TMB) substrate is used to show the color. Absorbance of the sample is negatively related to the β-Estradiol residual in it, after comparing with the Standard Curve, multiplied by the dilution multiples, β-Estradiol residual quantity in the sample can be calculated.

3. Applications

This kit can be used in quantitative and qualitative analysis of β-Estradiol residual in animal tissue (muscle, liver, fish and shrimp) feed and urine, etc.

4. Cross-reactions

β-Estradiol………………………………………100%

3–methylether–Estradiol ………………………30%

ethinylestradiol………………………………….7%

Estrone…………………………………..….…<0.1%

oestradiol benzoate ……………………..…<0.1%

estriol ………………………………………<0.1%

5. Materials Required

5.1 Equipments:

┅┅Microtiter plate spectrophotometer (450nm/630nm)

┅┅Rotary evaporator or nitrogen drying instruments

┅┅Homogenizer or stamocher

┅┅Shaker

┅┅Gyroscope  or vortex mixer

┅┅Centrifuge

┅┅Analytical balance (inductance: 0.01g)

┅┅Graduated pipette: 10ml

┅┅Rubber pipette bulb

┅┅Micropipettes: 20(l～200(l、100(l～1000(l

250(l-multipipette

Reagents：

----chloroform (AR)

---- Acetonitrile (AR) (for tissue samples)

----Sodium hydroxide (AR) (for tissue samples)

┅┅Concentrated phosphoric acid (the content is 85%) (AR)(For feed samples)

----Glucuronidase/ Arylsulfatase

----deionized water

6. Kit Components

1、Microtiter plate with 96 wells coated with coupling antigen

2、standard solutions(6 bottles:1ml/bottle)

0ppb，0.5ppb，1.5ppb，4.5ppb，13.5ppb，40.5ppb

3、High concentration standard control: (1ml/bottle) 1ppm

4、Enzyme-labeled secondary antibody solution (12ml)…red cap

5、Antibody solution ( 7ml) ……………………green cap

6、Solution A (7ml)……………………………white cap

7、Solution B (7ml)…………………………red cap

8、Stop solution (7ml) …………………………yellow cap

9、20×concentrated Wash solution(40ml) ………transparent cap

10、2×Concentrated extraction solution (50ml)………blue cap

7. Reagents Preparation:

Solution 1: 0.1 M sodium hydroxide solutions

Weigh 0.4g sodium hydroxide and dissolve with deionized water, dilute to 100ml;

Solution 2: 1M sodium hydroxide solution (for tissue sample)

Weigh 4g sodium hydroxide and dissolve with deionized water, dilute to 100ml;

Solution 3: acetonitrile-0.1M sodium hydroxide solution (for tissue sample)

Mix  84ml acetonitrile and 20ml 0.1M 氢sodium hydroxide solution, mix completely;

Solution 4: 6M phosphoric acid solution

Dilute 100ml phosphoric acid solution (the content is 85%) with deionized water, and dilute to 150ml;

solution 5: acetonitrile－chloroform solution

Mix 50ml acetonitrile and 100m chloroform, mix completely;

Solution 6: Extraction solution

Dilute the concentrated extraction solution with deionized water in the volume ratio of 1: 1, which will be used for sample extraction. This solution can be conserved for a month at 4℃.

Solution 7: Wash solution

Dilute the concentrated wash solution with deionized water in the volume ratio of 1:19(or according to the requirement), which will be used to rinse the plates. The diluted wash solution can be conserved for a month at 4℃.

8. Sample Preparations

8.1 Notice and precautions for the users before operation:

(a) Please use one-off tips in the process of experiment, and change the tips when absorbing different reagent.

(b) Make sure that all experimental instruments are clean.

(c) Keep untreated samples in freeze;

(d) Treated samples should be tested immediately, and they can not be conserved for over 4h at room temperature;

8.2 Tissue sample (chicken, chicken liver, pork, pig liver, fish and shrimp)

----Homogenize the samples with homogenizer or stamocher;

----Weigh 2.0±0.05g of the homogenate into a 50ml polystyrene centrifuge tube, then add acetonitrile-0.1M sodium hydroxide solution, shake for 10min with shaker, centrifuge at room temperature(20-25℃) for 10min, at least 3000g;

----Take 1ml of the supernate into a 10ml clean glass tube, dry with 50～60℃ water bath nitrogen gas flow;

----Dissolve the dry leftover with 0.5ml chloroform;

----Add 2ml 1M sodium hydroxide solution for reverse extraction, then shake fiercely for 5min, centrifuge at room temperature(20-25℃) for 10min, at least 3000g, take 1ml of the supernate and add 100(l 6M phosphoric acid solution ,mix completely;

----Add 5ml acetonitrile－chloroform solution, shake for 10min with shaker, centrifuge at room temperature(20-25℃) for 10min, at least 3000g, transfer the substrate phase into a 10ml clean glass tube, dry with 50～60℃ water bath nitrogen gas flow;

----Dissolve the dry leftover with 1ml extraction solution;

----Take 50(l of the prepared solution for assay;

8.3 Urine samples:

----Take 2ml urine sample into a centrifuge tube, centrifuge at room temperature(20-25℃) for 10min, at least 3000g, until the urine is transparent ;

----Transfer 1ml of the supernate into a 50ml polystyrene centrifuge tube, and then add 10(l Glucuronidase/ Arylsulfatase, hydrolyze for 3h at 37℃;

----Add 5ml chloroform and shake for 10min, then centrifuge at room temperature (20-25℃) for 10min, at least 3000g;

----Take 1ml of the substrate organic phase into a 10ml clean glass test tube, dry with 50～60℃ water bath nitrogen gas flow;

----Dissolve the dry leftover with 1ml extraction solution;

----Take 50(l of the prepared solution for assay;

8.4 Feed sample:

----Weigh 2.0±0.05g feed sample, then add 8ml acetonitrile, shake for 10min with shaker, then centrifuge at 15℃ for 10min, at least 3000g;

----Take 2ml of the supernate to a 10ml clean glass tube, dry with 50～60℃ water bath nitrogen gas flow;

----Then add 0.5ml chloroform , whorl for 20s with gyroscope  or vortex mixer, add 2ml 1M sodium hydroxide solution, whorl for 30s, then centrifuge at 15℃ for  5min, at least 3000g;

----Take 1ml of the supernate; add 100ml phosphoric acid solution, whorl for 10s with gyroscope  or vortex mixer;

----Dilute the extracted sample solution with extraction solution in the volume ration of 1:39(25(l extracted sample solution + 975(l extraction solution), then whorl to mix completely;

----Take 50(l of the water phase for assay;

Note: If the ration of the sample concentration against the dilution multiple is bigger than 40.5, please improve the dilution multiple (e.g. set the dilution multiple to 320), otherwise the test result will be beyond the scale of the calibration curve. Remember to use the real sample dilution multiple when processing data with the analytical software.

9. Assay process

9.1 Notice before assay:

9.1.1 Make sure all reagents and microwells are all at room temperature (20-25℃).

9.1.2 Return all the rest reagents to 2～8℃ immediately after used.

9.1.3 Washing the microwells correctly is an important step in the process of assay; it is the vital factor to the reproducibility of the ELISA analysis.

9.1.4 Avoid the light and cover the microwells during incubation.

9.2 Assay Steps:

9.2.1 Take all reagents out at room temperature (20-25℃) for more than 30min, homogenize before use.

9.2.2 Get the microwells needed out and return the rest into the zip-lock bag at 2-8℃ immediately.

9.2.3 The diluted wash solution should be rewarmed to be at room temperature before use.

9.2.4 Number: Numbered every microwell positions and all standards and samples should be run in duplicate. Record the standards and samples positions.

9.2.5 Add standard solution/sample: Add 50 µl of standard solution or prepared sample to corresponding wells. Add 50µl antibody solution. Mix gently by rocking the plate manually and incubate for 30min at 25℃ with cover.

9.2.6 Wash: Remove the cover gently and pure the liquid out of the wells and rinse the microwells with 250µl diluted wash solution (solution 7) at interval of 10s for 4~5 times. Absorb the residual water with absorbent paper (the rest air bubble can be eliminated with unused tip).

9.2.7 Enzyme-labeled secondary antibody: Add Enzyme-labeled secondary antibody solution 100(l to each well, Mix gently by rocking the plate manually and incubate for 30min at 25℃ with cover. Then repeat the wash step.

9.2.8 Coloration: Add 50µl solution A and 50µl solution B to each well. Mix gently by rocking the plate manually and incubate for 15min at 25℃ with cover(see 12.8).

9.2.9 Measure: Add 50µl the stop solution to each well. Mix gently by rocking the plate manually and measure the absorbance at 450nm against an air blank (It’s suggested measure with the dual-wavelength of 450/630nm. Read the result within 5min after addition of stop solution. ) (We can also measure by sight without stop solution in short of the ELIASA instrument)

10. Results

There are 2 different methods to determinate the results. Method 1 leads to a round estimation, and method 2 leads to definite quantity estimation. (Please notice: the absorption is inversely proportional to the β-Estradiol concentration in the sample.)

10.1 Round estimation

We can get the range of different strengths from the compare of average absorption and standards by sight. For example, the absorption of sample 1 is 0.101, sample 2 is 0.532, and the absorptions of β-Estradiol standard solutions: 1.615 (0ppb); 0.532(0.5ppb); 0.861(1.5ppb); 0.416(4.5ppb); 0.144(13.5ppb); 0.044(40.5ppb). So we can say the strength of diluted sample 1 is between 13.5ppb and 40.5ppb; and diluted sample 2 between 1.5 ppb to 4.5 ppb. In order to obtain the β-Estradiol actually contained in a sample, the diluted sample results must be further multiplied by the corresponding dilution factor.

10.2 Definite quantity estimation

(1) The mean values of the absorbance values obtained for the standards and the samples are divided by the absorbance value of the first standard (zero standard ) and multiplied by 100%. The zero standard is thus made equal to 100% and the absorbance values are quoted in percentages.

B

Absorbance (%) = —— ×100%

B0

B ——absorbance standard (or sample)

B0 ——absorbance zero standard

(2) To draw a standard curve: Take the absorbance value of standards as y-axis, semi logarithmic of the concentration of the β-Estradiol standards solution (ppb) as x-axis.

--- The β-Estradiol concentration of each sample (ppb), which can be read from the calibration curve, is multiplied by the corresponding Dilution factor of each sample followed, and the actual concentration of sample is obtained.

Please notice:

For evaluation of the ELISA kits special software has been developed for exact and rapid analysis. The analysis software can be ordered on request.

Sample dilution multiple:

Tissue sample(muscle, liver , fish and shrimp)：10

Urine sample：5

Feed sample：160

11. Sensitivity, accuracy and precision

Test Sensitivity: 0.5ppb

Detection limit:

Chicken, chicken liver, pork, pig liver, fish and shrimp…5ppb

Feed…………………………………80ppb

Urine…………………………………2.5ppb

Accuracy:

Feed / Tissue……………………… 80±15%

Urine………………………………… 80±10%

Precision:

Variation coefficient of the ELISA kit is less than 10%.

12. Notice

12.1 The mean values of the absorbance values obtained for the standards and the samples will be reduced if the reagents and samples have not been regulated to room temperature (20-25℃).

12.2 Do not allow microwells to dry between steps to avoid unsuccessful reproducibility and operate the next step immediately after tap the microwells holder.

12.3. Homogenize each reagent before using.

12.4. Keep your skin away from the stop solution for it is the 2M H2SO4 solution.

12.5 Don’t use the kits out of date. Don’t exchange the reagents of different batches, or else it will drop the sensitivity.

12.6 Storage condition:

Keep the ELISA kits at 2-8℃, do not freeze. Seal rest microwell plates, Avoid straight sunlight during all incubations. Covering the microtiter plates is recommended.

12.7 Indications for the reagents going bad:

Substrate solution should be abandoned if it turns colors.

The reagents may be turn bad if the absorbance value (450/630nm) of the zero standard is less than 0.5(A450nm＜0.5).

12.8 The coloration reaction need 10-15min after the addition of solution A and solution B; But you can prolong the incubation time ranges from 20min to more if the color is too light to be determined., never exceed 30min,On the contrary, shorten the incubation time properly.

12.9 The optimal reaction temperature is 25℃. Higher or lower temperature will lead to the changes of sensitivity and absorbance values.

13. Storage condition and storage period

Storage condition: 2-8℃.

Storage period: 6 months.