Competitive Enzyme Immunoassay Kit for

Quantitative Analysis of Erythromycin

1. Background

Erythromycin is a macrolide antibiotic, which is applied as antibacterial and anti mycoplasma infective, as well as pig feed additives improving feed transformation efficiency and increasing body weight. The maximum residue limit of erythromycin is prescribed to 100(g/kg in Doc.235 of MOA.

This kit is a new product for drug residual detection based on ELISA technology, which only costs 1h in each operation and can considerably minimize operation errors and work intensity compared with instrumental analysis.

2. Test Principle

This kit is based on indirect-competitive ELISA technology. The microtiter wells are coated with coupling antigen. Erythromycin residue in the sample competes with the antigen coated on the microtiter plate for the antibody. After the addition of enzyme labeled anti-antibody, TMB substrate is used to show the color. Absorbance of the sample is negatively related to the Erythromycin reside in it, after comparing with the Standard Curve, multiplied by the dilution multiple, Erythromycin residue quantity in the sample can be calculated.

3. Applications

This kit can be used in quantitative and qualitative analysis of Erythromycin residue in animal tissue (muscle, liver).

4. Cross-reactions

Erythromycin………………………………………100%

Erythromycin Thiocyanate…………………………114%

Erythromycin Ethylsuccinate………………………67.4%

tylosin……………………………<0.1%

Tilmicosin…………………………<0.1%

spiramycin…………………………………<0.1%

Kitasamycin……………………<0.1%

5. Materials Required

5.1 Equipments:

┅┅Microtiter plate spectrophotometer (450nm/630nm)

┅┅Rotary evaporator or nitrogen drying instruments

┅┅Homogenizer or stamocher

┅┅Shaker

┅┅Gyroscope  or vortex mixer

┅┅Centrifuge

┅┅Analytical balance (inductance: 0.01g)

┅┅Graduated pipette: 10ml

┅┅Rubber pipette bulb

┅┅Volumetric flask: 500ml

┅┅glass bottle:10ml

┅┅Polystyrene Centrifuge tubes:2ml、50ml

┅┅Glass centrifuge tube: 10ml

┅┅Micropipettes:  20(l～200(l、200(l～1000(l

250ul-multipipette

5.2 Reagents:

┅┅Acetonitrile (AR)

┅┅Deionized water

6. Kit Components

1、Microtiter plate with 96 wells coated with coupling antigen

2、standard solutions(6 bottles×1ml/bottle)

0ppb, 0.2 ppb, 0.6 ppb, 1.8ppb, 5.4 ppb, 16.2ppb

3、High concentration standard control:(1ml/bottle)  1ppm

4.Enzyme-labeled secondary antibody solution(7ml)……red cap

5、antibody working fluid  7ml ……green cap

6、Solution A (7ml)……………………………white cap

7、Solution B (7ml)…………………………red cap

8、Stop solution (7ml) …………………………yellow cap

9、20×concentrated wash solution  40ml …transparent cap

10、2×concentrated extraction solution 50ml…… blue cap

7. Reagents Preparation:

Solution 1: extraction solution

Dilute the 2×concentrated extraction solution with deionized water in the volume ratio of 1:1, which will be used for sample extraction, this solution can be stored at 4℃ for 1 month.

Solution 2: wash solution

Dilute the 20×concentrated wash solution with deionized water in the volume ratio of 1:19, which will be used for washing the plates. This solution can be stored at 4℃ for 1 month.

8. Sample Preparations

8.1 Notice and precautions for the users before operation:

(a) Please use one-off tips in the process of experiment, and change the tips when absorb different reagent.

(b) Make sure that all experimental instruments are clean ,otherwise it will effect the assay result.

8.2 Tissue(muscle and liver):

┅┅Homogenizer or stamocher the sample with homogenizer or stamocher;

┅┅Weigh 2.0±0.05g homogenate into a 50ml polystyrene centrifuge tube; add 8ml Acetonitrile, mix completely for 10min.

┅┅centrifuge at least 3000g,at room temperature (20-25℃) for 10min.

┅┅Take 2ml of the supernate to a 10ml clean glass test tube, dry with 50～60℃ water bath nitrogen gas flow.

┅┅Dissolve the dry leftover with 2ml extraction solution(See solution 1).

┅┅Take 50(l of the prepared solution for assay.

9. Assay process

9.1 Notice before assay:

9.1.1 Make sure all reagents and microwells are all at room temperature (20-25℃).

9.1.2 Return all the rest reagents to 2～8℃ immediately after used.

9.1.3 Washing the microwells correctly is an important step in the process of assay; it is the vital factor to the reproducibility of the ELISA analysis.

9.1.4 Avoid the light and cover the microwells during incubation.

9.2 Assay Steps:

9.2.1 Take all reagents out at room temperature (20-25℃) for more than 30min, homogenize before use.

9.2.2 Get the microwells needed out and return the rest into the zip-lock bag at 2-8℃ immediately.

9.2.3 The diluted wash solution should be rewarmed to be at room temperature before use.

9.2.4 Number: Number every microwell position and all standards and samples should be run in duplicate. Record the standards and samples positions.

9.2.5 Add standard solution/sample: Add 50 µl of standard solution or prepared sample to corresponding wells. Add 50µl enzyme-labeled secondary antibody solution and 50µl antibody solution. Mix gently by shaking the plate manually and incubate for 40min at 37℃ with cover.

9.2.6 Wash: Remove the cover gently and pure the liquid out of the wells and rinse the microwells with 250µl diluted wash solution (See solution 2) at interval of 10s for 4~5 times. Absorb the residual water with absorbent paper (the rest air bubble can be eliminated with unused tip).

9.2.7 Coloration: Add 50µl solution A and 50µl solution B to each well. Mix gently by shaking the plate manually and incubate for 15 min at 37℃ with cover(see 12.8).

9.2.8 Measure: Add 50µl the stop solution to each well. Mix gently by shaking the plate manually and measure the absorbance at 450nm against an air blank (It’s suggested measure with the dual-wavelength of 450/630nm. Read the result within 5min after addition of stop solution. ) (We can also measure by sight without stop solution in short of the ELIASA instrument).

10. Results

There are 2 different methods to determinate the results. Method 1 leads to a round estimation, and method 2 leads to definite quantity estimation. (Please notice: the absorption is inversely proportional to the Erythromycin concentration in the sample.)

10.1 Round estimation

We can get the range of different strengths from the compare of average absorption and standards by sight. For example, the absorption of sample 1 is 0.6, sample 2 is 1.0, and the absorptions of Erythromycin standard solutions: 1.500(0ppb); 1.280(0.2ppb);0.950(0.6ppb);0.700(1.8ppb); 0.450(5.4ppb); 0.150(16.2ppb). So we can say the strength of diluted sample 1 is between 1.8ppb and 5.4ppb; and diluted sample 2 between 0.2ppb to 0.6ppb. In order to obtain the Erythromycin actually contained in a sample, the diluted sample results must be further multiplied by the corresponding dilution multiple.

10.2 Definite quantity estimation

(1) The mean values of the absorbance values obtained for the standards and the samples are divided by the absorbance value of the first standard (zero standard ) and multiplied by 100%. The zero standard is thus made equal to 100% and the absorbance values are quoted in percentages.

B

Absorbance (%) = —— ×100%

B0

B ——absorbance standard (or sample)

B0 ——absorbance zero standard

(2) To draw a standard curve: Take the absorbance value of standards as y-axis, semi logarithmic of the concentration of the Erythromycin standards solution (ppb) as x-axis.

--- The Erythromycin concentration of each sample (ppb), which can be read from the calibration curve, is multiplied by the corresponding dilution multiple of each sample followed, and the actual concentration of sample is obtained.

For evaluation of the ELISA kits special software has been developed for exact and rapid analysis. The analysis software can be ordered on request.

Dilution multiple:

Dilution multiple of samples:  4

11. Sensitivity, accuracy and precision

Test Sensitivity：0.2 ppb

Detection limit：

Tissue sample…………………………………0.8ppb

Accuracy:

muscle……………………………………………… 85±15%

liver ……………………………………………… 85±15%

Precision:

Variation coefficient of the ELISA kit is less than 10%.

12. Notice

12.1 The mean values of the absorbance values obtained for the standards and the samples will be reduced if the reagents and samples have not been regulated to room temperature (20-25℃).

12.2 Do not allow microwells to dry between steps to avoid unsuccessful reproducibility and operate the next step immediately after tap the microwells holder.

12.3. Homogenize each reagent before using.

12.4. Keep your skin away from the stop solution for it is the 2M H2SO4 solution.

12.5 Don’t use the kits out of date. Don’t exchange the reagents of different batches, for it will drop the sensitivity.

12.6 Storage condition:

Keep the ELISA kits at 2-8℃,do not freeze. Seal rest microwell plates Avoid straight sunlight for the standard sample and the colorless chromogenic reagent are sensitive to light.

12.7 Indications for the reagents going bad:

Substrate solution should be abandoned if it turns colors.

The reagents may be turn bad if the absorbance value (450/630nm) of the zero standard is less than 0.5(A450nm＜0.5).

12.8 The coloration reaction needs 15min after adding Solution A and Solution B. And you can prolong the incubation time ranges from 15min to 25min if the color is too light to be determined., never exceed 30min,On the contrary, shorten the incubation time properly.

12.9 The optimal reaction temperature is 37℃. Higher or lower temperature will lead to the changes of sensitivity and absorbance values.

13. Storage condition and storage period

Storage condition: 2-8℃.

Storage period: 6 months.