QuantiChromTM Peroxide Assay Kit (DIOX-250)

Quantitative Colorimetric Peroxide Determination at 585nm

DESCRIPTION

concentration is 100µM (labeled “Premix”). Dilute standard as follows.

concentration is 100µM (labeled “Premix”). Dilute standard as follows.

| No | Premix + H2O | Vol (µL) | H2O2 (µM) |
| --- | --- | --- | --- |
| 1 | 100µL + 0µL | 100 | 100 |
| 2 | 80µL + 20µL | 100 | 80 |
| 3 | 60µL + 40µL | 100 | 60 |
| 4 | 40µL + 60µL | 100 | 40 |
| 5 | 30µL + 70µL | 100 | 30 |
| 6 | 20µL + 80µL | 100 | 20 |
| 7 | 10µL + 90µL | 100 | 10 |
| 8 | 0µL + 100µL | 100 | 0 |

| No | Premix + H2O | Vol (µL) | H2O2 (µM) |
| --- | --- | --- | --- |
| 1 | 100µL + 0µL | 100 | 100 |
| 2 | 80µL + 20µL | 100 | 80 |
| 3 | 60µL + 40µL | 100 | 60 |
| 4 | 40µL + 60µL | 100 | 40 |
| 5 | 30µL + 70µL | 100 | 30 |
| 6 | 20µL + 80µL | 100 | 20 |
| 7 | 10µL + 90µL | 100 | 10 |
| 8 | 0µL + 100µL | 100 | 0 |

2.	Transfer 20 µL diluted standards in duplicate and each sample in quadruplicate into wells of a clear bottom 96-well plate. Add 100

µL Working Reagent WRa to all standards and to two wells of the sample, add 100 µL control WRb to the other two sample wells.

3.	Incubate 30 min at room temperature and read optical density at 540-610nm (peak absorbance at 585nm).

CALCULATION

Subtract blank OD (water, #8) from the standard OD values and plot the OD against H2O2 concentrations (see Calibration curve). For samples, subtract background OD measured with WRb from sample OD measured with WRa. Calculate the sample peroxide content by comparing the sample absorbance ÄOD and the calibration curve.

Conversions: 1 µM H2O2 equals 34 ng/mL or 34 ppb.

MATERIALS REQUIRED, BUT NOT PROVIDED

Pipeting devices and accessories, 96-well plates and plate reader.

EXAMPLES:

Duplicate assays for goat serum, human serum, 293 cell culture medium and fresh human urine gave peroxide content of 8.7± 2.8, 14.2 ± 2.7, 2.4 ± 0.0 and 1.4 ± 0.6 µM (n = 2).

2.	Transfer 20 µL diluted standards in duplicate and each sample in quadruplicate into wells of a clear bottom 96-well plate. Add 100

µL Working Reagent WRa to all standards and to two wells of the sample, add 100 µL control WRb to the other two sample wells.

3.	Incubate 30 min at room temperature and read optical density at 540-610nm (peak absorbance at 585nm).

CALCULATION

Subtract blank OD (water, #8) from the standard OD values and plot the OD against H2O2 concentrations (see Calibration curve). For samples, subtract background OD measured with WRb from sample OD measured with WRa. Calculate the sample peroxide content by comparing the sample absorbance ÄOD and the calibration curve.

Conversions: 1 µM H2O2 equals 34 ng/mL or 34 ppb.

MATERIALS REQUIRED, BUT NOT PROVIDED

Pipeting devices and accessories, 96-well plates and plate reader.

EXAMPLES:

Duplicate assays for goat serum, human serum, 293 cell culture medium and fresh human urine gave peroxide content of 8.7± 2.8, 14.2 ± 2.7, 2.4 ± 0.0 and 1.4 ± 0.6 µM (n = 2).

0.6

0.6

0.4

0.4

0.2

0.2

0

0

Peroxide

Peroxide

1.2

1.2

1

1

0.8

0.8

Peroxide (e.g. hydrogen peroxide H2O2) is one of the key reactive oxygen species formed under oxidative stress conditions. High levels of peroxide formation have been linked to pathological conditions such as ageing, asthma, diabetes, atherosclerosis, cataract, inflammatory arthritis and neurodegenerative diseases.

Simple, direct and automation-ready procedures for quantitative determination of peroxide find wide applications in research and drug discovery. BioAssay Systems' peroxide assay kit is designed to measure peroxide concentration in biological samples without any pretreatment. The improved method utilizes the chromogenic Fe3+-xylenol orange reaction, in which a purple complex is formed when Fe2+ provided in the reagent is oxidized to Fe3+ by peroxides present in the sample. The intensity of the color, measured at 540-610nm, is an accurate measure of the peroxide level in the sample. The optimized formulation substantially reduces interference by substances in the raw samples.

KEY FEATURES

Sensitive and accurate. Enhanced color intensity using sorbitol. Detection range 0.4 µM (14 ng/mL) to 100 µM (3,400 ng/mL) H2O2 in 96-well plate assay.

Simple and high-throughput. The procedure involves addition of a single working reagent and incubation for 30 min. Can be readily automated as a high-throughput assay in 96-well plates for thousands of samples per day.

Potential interference reduced. For each sample, a background control is run under identical conditions with the omission of the Fe2+. This procedure corrects for any interference from endogenous iron present in biological samples.

APPLICATIONS:

Direct Assays: H2O2 in biological samples (e.g. serum, citrate-plasma, urine, cell lysate, culture medium).

Pharmacology: effects of drugs on peroxide metabolism.

KIT CONTENTS (250 tests in 96-well plates)

Reagent A: 1 mL	Reagent B (for background control): 1 mL

Reagent C: 50 mL

H2O2 Standard: 1 mL 3% stabilized H2O2.

Kit shipped at room temperature.

Storage conditions. The kit is shipped at room temperature. Store all reagents at 4 °C. Shelf life of at least 6 months (see expiry dates on labels).

Precautions: reagents are for research use only. Normal precautions for laboratory reagents should be exercised while using the reagents. Please refer to Material Safety Data Sheet for detailed information.

PROCEDURES

Sample Treatment: several chemicals are known to interfere and should be avoided in sample preparation. These include ascorbic acid, EDTA, heparin, DMSO (>0.02%), NP-40 (>0.6%), SDS (>0.12%), Tris (>8mM) and ethanol (>0.4%).

0	20	40	60	80	100

[H2O2], µM

Standard Curve in 96-well plate assay

LITERATURE

1.	Karageuzyan KG (2005). Oxidative stress in the molecular mechanism of pathogenesis at different diseased states of organism in clinics and experiment. Curr Drug Targets Inflamm Allergy. 4(1):85-98.

2.	Arab K, Steghens JP. (2004). Plasma lipid hydroperoxides measurement by an automated xylenol orange method. Anal Biochem. 325(1):158-63.

3.	Banerjee D, Madhusoodanan UK, Nayak S, Jacob J. (2003). Clin Chim Acta. Urinary hydrogen peroxide: a probable marker of oxidative stress in malignancy. 334(1-2): 205-9.

0	20	40	60	80	100

[H2O2], µM

Standard Curve in 96-well plate assay

LITERATURE

1.	Karageuzyan KG (2005). Oxidative stress in the molecular mechanism of pathogenesis at different diseased states of organism in clinics and experiment. Curr Drug Targets Inflamm Allergy. 4(1):85-98.

2.	Arab K, Steghens JP. (2004). Plasma lipid hydroperoxides measurement by an automated xylenol orange method. Anal Biochem. 325(1):158-63.

3.	Banerjee D, Madhusoodanan UK, Nayak S, Jacob J. (2003). Clin Chim Acta. Urinary hydrogen peroxide: a probable marker of oxidative stress in malignancy. 334(1-2): 205-9.

Samples can be analyzed immediately after collection, or stored in aliquots at –20 °C. Avoid repeated freeze-thaw cycles. If particulates are present, centrifuge sample and use clear supernatant for assay.

Reagent Preparation: prepare at least 100 µL/well Working Reagent (WRa) by mixing 1 volume of Reagent A with 100 volumes Reagent C. Prepare the control reagent (WRb) by mixing 1 volume of Reagent B with 100 volumes Reagent C. Equilibrate to room temperature before assay.

Procedure using 96-well plate:

1. Standards. Prepare fresh standards on the day of assay. Pipette 10 µL 3% H2O2 and mix well with 990 µL H2O in a 1.5-mL Eppendorf tube. Mix 10 µL of this solution with 872 µL H2O. The final H2O2