Favor Cool-Heating Dry Bath Incubator

Favor Cool-Heating Dry Bath Incubators accommodate an assortment of interchangeable Heating Block Modules. These reliable and economical products are ideal for a variety of applications including restriction enzyme of DNA digestion, DNA denaturation, BUN, melting agar, coagulation studies, in situ hybridization and Hot Start PCR. Favor Cool- Heating Dry Bath Incubators have one-year warranty.

Favor Cool- Heating Dry Bath Incubators are available in one block capacity models. These reliable, economical and convenient incubators are equipped with multiple specifications of block modules to serve as ideal instruments for virtually every conceivable application.

Features

- Microprocessor controller with digital performance

- The block designed by tube shape is easy to heat uniformly

- Fast Heating by P.I.D. controller

- Thermostat temperature from -10~ 100 ℃

- The fidelity of temperature range is 0.1 ℃

- LCD monitor

- User Calibration

- Over heating protection device ensures safety & reliability.

- Beep-signal and stop after program completion

Experiment Information of Block from Favorgen

Experiment Information of Block from Favorgen

| Type of Block | Parameter | Setting Temperature | Tube Temperature (I), time waste (I) | Tube Temperature (II), time waste (II) |
| --- | --- | --- | --- | --- |
| FATSM-B | 54X0.5ml | 50℃ | 49.6℃, 90 seconds | 49.8℃, 120 seconds |
| FATSM-C | 35X1.5ml | 50℃ | 49.3℃, 120 seconds | 49.6℃, 150 seconds |

Specification

| Controller | Digital Microprocessor controller |
| --- | --- |
| Display | 20 x 2 LCD |
| Temperature Control Range | -10~100℃ in the environment of 25℃ |
| Temperature increase | 0.1℃ |
| Temperature Calibration | Yes |
| Temperature uniformity | ±0.3℃ |
| Temperature accuracy | ±0.3℃ |
| Operation mod | a. Constant operation: Constant temperature(-5~110℃) Timer: 1min~ 99hours 59 mins |
|  | b. Program operation: Programmable: 1-2 steps Timer: 1min-99hours 59 mins for each step |
| Data Log | No |
| Operating Temperature | 5~30℃ |
| Safety Device | Leakage proof for heating chamber |
| Heating Chamber Material | Plastics |
| Block Material | Aluminum |
| Block Dimension | A Block=L136 x W126 x D28 mm B Block= L136 x W126 x D33 mm C or D Block= L136 x W126 x D44 mm |
| Block Type | Standard and customized types are available Four types of Standard block: A-block: 96 x 0.2ml B-block: 54 x 0.5ml C-block: 35 x 1.5ml D-block: 35 x 2.0ml |
| Rated Voltages | 110V/220V selectable |
| Unit Dimension | W212x L300 x H180 mm |
| Weight | 4.6 kg |

Block Type

| Cat. No. | Product Name | Descriptions | Max. Temp | Min Temp. | Max.Speed (rpm) |
| --- | --- | --- | --- | --- | --- |
| FATSM-A |  | 96×0.2ml | 100℃ | -10℃ | 1500 |
| FATSM-B |  | 54×0.5 ml | 100℃ | -10℃ | 1500 |
| FATSM-C |  | 35×1.5ml | 100℃ | -10℃ | 1500 |
| FATSM-D |  | 35×2.0ml | 100℃ | -10℃ | 1500 |
| FATSM-E |  | 15×0.5 ml+20×1.5ml | 100℃ | -10℃ | 1500 |
| FATSM-F |  | 24×Dia.≤φ12mm tube | 100℃ | 0℃ | 1200 |
| FATSM 02 |  | 12×15ml Falcon | 100℃ | 0℃ | 750 |
| FATSM 05 |  | 6×50ml Falcon | 100℃ | 0℃ | 750 |

Ordering Information

| Catalog # | Product Name | Block Capacity | MaxTemp | Dimensions (LxWxH) |
| --- | --- | --- | --- | --- |
| FATRB 003 | Favor Cool-Heating Dry Bath Incubator | 1 Block | 100ºC | 300mm x 212mm x 180mm |